## Slide 1: Shielding a neutron source : a tough job but possible
ISNIE Summer school PSI, ch 20/9/2018
A.Menelle
LLB CEA-CNRS Saclay

## Slide 2: Teoretical OUTLINE
What Sylvain Desert ask me to do
Part 1 Theory
  Types of radiation
  Radiation from different sources
  How shielding works
  Shielding materials (for different uses)
  Thickness approximation (Sullivan hand calcs)
  Leak paths
  Activation
Part 2 Engineering consideration
  Cost
  Structural issues (stability/seismic/floor load)
  Disposal
  Construction techniques
  Ease of manufacture
Part 3 Shielding (Low energy) ie out of line of sight (similar reactor – spallation)
  Instrument caves
  Beamline shielding
  Use of materials (cost/performance/ease of use/disposal)
  Applicable Construction techniques (cast in place/blocks)
  Worked example
    Reactor source
    Spallation source

## Slide 3: Pratical OUTLINE
What I will do
Part 1 Theory
  Why a shielding ?
  Workers protection regulations
  Blue green yellow orange and red zones
  The different radiations
  Activation example
  Neutron specificity
  Detector protections
Part 2 Engineering consideration
  Basic formulae
  Cross sections issues
  Materials to use
  Transmission / absorption measurements
Part 3 Shielding
  Gamma shielding
  Neutron absorption
  Fast neutron absorption
Part 4 Various
  Other common materials

## Slide 4: Why Shielding : workers and detectors cases
Two different jobs for shielding
Shielding for people : enable them to work within authorized dose limits
Shielding for detectors : detect only neutrons scattered from the sample
People protection parameters
Distance
Time of work
Shielding
Classification of people
Detectors protection parameters
Absoption of unwanted particules
Efficiency of detection of unwanted particules by the detectors

## Slide 5: Protection of workers : dose limits
May depend a bit on local regulations, though quite homogeneous through europe.
In France, 3 categories of workers on installations with ionizing radiations and their dose limits :
A : 20 mSv/year
B : 6 mSv/year (CEA : <200 mSv/day)
Non classified worker (NE): 1 mSv/year (CEA : 1 mSv/day and dose rates > 10 mSv/h not allowed)
(In a monochromatic beam on the sample in Orphée : 5 Sv/h)
Deal with 2 types of workers
Staff of the installation : B (to give them a specific medical survey)
Users : NE

## Slide 6: Protection of workers : zones
Access control and limits will depend on the radiation doses present in the different areas of the buildings
Design buildings and shieldings so that users may have access to what they need as non classsified workers (NE) :
Experiment cabin and cave, pathways and preparation laboratories : blue < 7.5 mSv/h (~200 n/cm2/s)
Equipment necessary for the source operation : green < 25 mSv/h
Zone | Limits | Access
Non exposed | < 0.5 mSv/h | No limitation of access
Blue | < 7.5 mSv/h | Controlled access; no limitations
Green | < 25 mSv/h | Controlled acces with restrictions for NE
Yellow | < 2 mSv/h | Regulated access, no NE
Orange | < 100 mSv/h | Access submitted to authorizations
Red | > 100 mSv/h | Forbidden access

## Slide 7: Different type of radiations
Neutrons
Low interaction with matter, large mean free path, energetic ones difficult to stop
Bounce on atoms before being absorbed  random walk : thermal neutrons are like a gas; beware of leaks
Absorption creates activation  with 25 meV you may obtain 1MeV at a different place and time
To calculate exposition and dose rates, sum up all radiations
Particles that mainly go straight and just transfer their energy
Alpha (4He); stopped by a sheet of paper. Danger in case of contamination and internal exposure. Do not allow to drink or eat; access prohibited with open wound; control radioactivity in air.
Beta (e-); stopped by a few mm of any material (metal, glass, plastic, …)
Gamma high energy photon; needs mass to be stopped (lead, tungsten, water, …). ! At high energy, may break a nucleus and expel neutrons !

## Slide 8: Activation example
Air contains 0.9% of 40Ar. Under thermal neutron irradiation, you create radioactive 41Ar with half life 1.8 hour
In Orphée : part of beam of 108 n/cm2/s in air of the 14 instruments create around 105 Bq/h of 41Ar. If confined and breathed by a user during 15mn, it might create an internal exposure of 1 mSv
ESS Magic : 109 n/cm2/s on sample, create 106 Bq/h, might induce a dose of 10 mSv in 15mn by breathing. This is 10 time the day dose limit for NE in CEA.
Solutions
Avoid air irradiation; put beam under vacuum
Vent the experimental caves
Be prepared to have questions of the safety authorities

## Slide 9: Neutrons specificity
Fission fast neutrons (> 1eV)
Low interaction, low detection, cross everything, no capture, no creation of radio-isotope, low energy deposition, go straight
But
High slowing down interaction with light atoms and goes to thermal neutrons : beginning of pb
Thermal neutrons (< 1eV)
Lots of interactions with atoms, multi bounces, a gaz of neutrons; no leaks is a challenge
lots of activation channels of nearly all atoms, espeacially the heavy ones
To avoid pb : choose carrefully materials to be used and pay a lot of attention to additives or impurities !!!!!!
1 mg of Co in 1g of Si irradiated 1 day with 1013n/cm2/s. 1 month later : 0 Bq of Si, but 1352 Bq of 60Co with 5.3 year half life. Think of the Pb with 1 ton of shielding !!!!!!!

## Slide 10: Shielding for detectors
Long and time consuming efforts to decrease instrument background
My targets :
source on, beam off : 10 times the electronic background
source on, beam on : 100 times the electronic background
Obtained by protection against outside neutrons : shielding of the cave and the detector vessel with neutron absorbers
Special attention for fast neutrons : low efficiency of detection (< 10-6 for 1 MeV fast neutrons) but always there.
Thermalization with concrete or PEHD
Absorption with Cd of B4C
Beware of the holes !!!
          Think if it is useful before thermalizing

## Slide 11: Shielding for fast neutron
Not to do !
Fast neutrons present between guides 4 and 5
You think that neutrons comes from the reactor
Install a wall of PEHD / B4C to thermalized and absorb them
No change !!!
Neutron dose
before adding a wall
4 mSv/h fast
4 mSv/h thermal
Neutron dose
with the wall
4 mSv/h fast
4 mSv/h thermal

## Slide 12: Shielding for fast neutron
To do !
Put absorber (concrete) everywhere !

## Slide 13: Shielding for detectors : Theory
Absorbs neutrons coming from outside
Cover the detector vessel and sample vessel with neutron absorbers
Add PEHD outside if too much* fast neutrons around
Detectors must see only the sample
Use collimators in front of the detector (fixed or moving) : Gd, Cd, B
Put nothing except sample in the neutron beam (windows create neutron scattering)
* : too much depends on the spectrometer. In Orphée, high fast neutron level betwwen G4 and G5 (2.5 mSv/h). Pb for reflectometer, not for powder diffractometer.

## Slide 14: Radial Collimators and beam setup
5C1 collimator
Shield flight tank with neutron absorber
Avoid anything in the beam in view of detectors

## Slide 15: Shielding for detectors :optimization
A long work
Counts/hour/ time channel on a reflectometer
Beam on with sample
Beam off
Different
conditions
Source off
Electronic noise

## Slide 16: What to use and where ?
Materials for spectrometers
Materials for shieldings
Special materials

## Slide 17: Basic formulae
Transmission through a thickness x :
x
t : transmission (absorption a = 1-t )
 : cross section
 : density
x : thickness
For solids :
 : cross section in barns
 : density in g/cm3
x : thickness in cm
M : molar mass in g
Example : 1cm Al
M=27g; =0.23barns; =2.7g/cm3
t = 0.986 gives a=0.014
For gases :
 : cros section in barns
x : thickness in cm
P : pressure in bars
Example : 5 bars 3He on 1cm
: 5333 barns
t = 0.49 gives a=0.51

## Slide 18: Which cross section ?
a : absorption cross section : prop. to l
c : coherent scattering cross section : constant with l for l<lc
At least a Bragg peak must be there for single crystals
i : incoherent scattering cross section : constant
lc given by 2d sin(q/2)=lc with q=180° scattering angle
for d=2.5Å (max distance between atomic planes) it gives lc=5Å
Example : Al
a = 0.23 barns c = 1.495 barns i = 0.008 barns
Beware :
Cross sections at 2200 m/s in tables (1.8 Å thermal neutrons)
Correction a(l) : l* a /1.8
l<lc
l<lc
l>lc

## Slide 19: Cross sections variations
Cd
Gd
Not always prop. to l
Low energy
High energy

## Slide 20: Depends of the microstructure
This does not take into account heterogeneities small angle scattering.
Example : porous alumina made by K.Lagrene :
Al2O3
a = 0.46 barns c = 15.68 barns i = 0.02 barns
Should be transparent to neutrons
But not at all due to SANS of pores.
Might be reduced if you fill pores with contrast match solvent.

## Slide 21: Common materials used on spectrometers
Al (AG3, AG3-net or AG5)
  Light, easy to machine
  Transparent with neutrons
  Low activation (espeacilly for AG3-net), half life Al : 2.24 mn
  Always wait at least 10mn before any Al sample holder manipulation after irradiation: emission of a g of 0.8 MeV
Steel
  Try to use steel without Co (it activates with a half life of 5.3 year, g of 1.3 MeV
  Quite absorbing a = 2.56 barns
  Strong
Zircalloy (Zr + 2.5% Sn)
  Transparent to neutrons
  High strength, low activation, (emits béta)
  Difficult to find

## Slide 22: Other materials
Brass (Cu Zn alloy)
  Easy to machine but
  Zn has high activation; half life 244 days, g of 1.3 MeV
  To avoid close to the beam
Plastics
  Becomes powder under g irradiation. g breaks double links in the molecules
  Highly absorbing because of hydrogen : i = 80 barns
  To avoid if high g dose

## Slide 23: Various thermal neutrons Transmissions
Fe
0
1
4Å
18Å
H2O
D2O
1
0
4Å
18Å
Transmission of Zircalloy windows with thickness and wavelength

## Slide 24: Gamma radiation protection
Concrete
  Contain roughly 15% of water (for fast neutrons thermalization)
  Low cost, easy to handle
  Various quality available
    Normal concrete : 2.3 kg/dm3
    Concrete with baryty : 3.6 kg/dm3
    Heavy concrete : usual 4.4 kg/dm3, special 6 kg/dm3, for 2T 7 kg/dm3
  Add absorbing materials for neutrons (B203) or g (Fe, Pb, …)
  May be a bit like a cooking receipe (heavy concrete for 2T)
Lead
  Not that expensive, easy to handle and design, available as bricks
  Density : 11.3 kg/dm3
  Beware : transparent to thermal neutrons a = 0.2 barns
Tungsten
  Highly absorbing; Density : 19.3 kg/dm3; a = 18.3 barns
  Expensive and difficult to machine
To be protected against g, you have to put mass !

## Slide 25: Gamma absorbing power

## Slide 26: Materials for neutron absorption
Most commonly used :
Li ; a = 70 barns (6Li ; a = 940 barns , !! nuc. mat !!)
B ; a = 767 barns (10B ; a = 3835 barns)
Cd ; a = 2520 barns
Gd ; a = 49700 barns; qq mm are suffisent !
Beware :
Neutron absorption conduct generally to the emission of a high energy particle; usually a g, sometime a a or a b.
The heavier the absorbing nucleus, the highest number of desexcitation paths with high energy particle emission
High absorption neutron cross section.

## Slide 27: Neutron absorption
Cd (metallic foils)
  Emits g mainly at 0.6 MeV, but also at 6 MeV
  Easy to handle and use
  Need a limited thickness
  Few neutron reflection
  Regulations for handling more and more stric (poison)
B4C (powder, rubber, plastic, sintered, carbon)
  Difficult to machine, but may be molded and cut
  Emits g of 0.5 MeV only
  Reflects a bit the neutrons
Boron nitrides (white friable blocs)
  Special for furnaces
  Easy to machine before heating
  Reflects a lot of neutrons
6Li (plastic LiF foil or carbonate of Li powder)
  Absorbs without g emission, but emits a a of 4.5 MeV
  Very expensive, easy to handle, use only in special cases
Gd
  Very efficient absorber
  Emits energetic g around 0.2 MeV, 1 Mev and 6.5 MeV (use with moderation)
Boral plates (B4C between Al fols)
  Very efficient absorber
  Difficult to machine
  Difficult to find

## Slide 28: Neutron absorption : one more fast neutrons Pb
Be carrefull : B and Li emit fast neutrons around 2MeV.
  Efficiency for B : 0.5 10-6
  Efficiency for Li : 80 10-6
Be carrefull : high energy g can extract neutrons from heavy nuclei and can emit fast neutrons
Example : neutrons/Cd/Pb or neutrons/Gd/Pb
  Low efficiency : ~ 10-8
Detection efficiency of fast neutrons by 3He gaz detectors
  a = 0.8 barns; detection efficiency 6 10-4
  White beam (109n/s) absorbed by Li, emits 8 104n/s
  30 cm2 neutron detector at 1 m sees 600 n/s (solid angle)
  Low efficency of detection  background of 0.4 n/s
  You cannot avoid shielding of fast neutrons !!

## Slide 29: Stop fast neutrons
Stop fast neutrons with thermalization !!
Best thermalizer is H atom.
Various solutions tested :
Water tank : leaks Pb
PEHD (CH2)n (easy to handle but will not sustain high flux)
Concrete (15% water inside)
!! Do not forget to put an absorber of thermal neutrons afterwards !!
Slowing down from 2MeV to 0.025eV | H2O | D2O
Mean free path (cm) | 0.5 | 2.5
Slowing down distance (cm) | 5.7 | 11
Nb chocs | 18 | 25

## Slide 30: Neutron absorbers and background
All neutrons absorbers reflects more or less thermal neutrons
15mn on G61 | Diffus | Peak
6Li | 5 | 5
Cd | 20 | 170
B4C sintered | 5 | 700
B4C carbon | 40 | 1800
B4C plastic | 590 | 1100
Boron nitride | 50 | 10000
6Li
Cd
B4C plastic

## Slide 31: Materials with specific use 1/2
Si
  Large size and low price single crystals (20*20*30 cm3).
  Transparent to neutrons.
  May be used at low and high T°
  A bit fragile
Quartz (Si02)
  Transparent to neutrons and light
  Available on different size
  More expensive than Si
Glasses
  Composition not always known with precision
  Often highly absorbing
  Schott glass B270 quite transparent
  95% transmission after 4mm.

## Slide 32: Materials with specific use 2/2
.
Alumina (Al2O3)
  Available as single crystal
  Transparent for neutrons
  May be used at high and low T°
Be
  Short wavelength absorber
HOPG graphite filters
  Selective wavelength absorption
Fast neutrons Bi filter
  High mass to stop fast neutrons (209)
  low absorption : a = 0.04 barns

## Slide 33: Conclusion
Lots of work and tests to be done
Thank you

## Slide 34: Glasses for neutron guides
Compositions des verres à guides de neutrons en % en masse
Float Borkron A8866 Borofloat
SiO2 70.62 62 67 83
Al2O3 0.9 0 0.9 2.7
ZnO2 0 11 0 0
Na2O 13.6 7 0.5 3.6
B2O3 0 12.5 19 12.9
CaO 9.8 0 0 0
MgO 3.85 0 0 0
K2O 0.5 0 8.3 0.6
Fe2O3 0.1 0 0 0.001
SO3 0.28 0 0 0
TiO2 0.015 0 0 0
LiO2 0 0 0.9 0
As2O3 0 0 0.2 0
ZrO2 0 0 0 0.05
99.665 92.5 96.8 102.851
Rôles :
Guider les neutrons
Tenir le vide
Absorber les neutrons
Float :
Crée des g
Peu cher
Surface de qualité
Un peu mou
Pas d’absorption neutrons
Borkron (guides du LLB):
Cher
A polir
Résistant
Borofloat :
Peu cher
Surface de qualité
Peu résistant au rayonnement
A8866 :
Alternative au Borkron, mais trop dur (difficile à polir)